Experimenter’s RF Spectrum Analyzer

George R. Steber, WB9LVI

Authors notes on spectrum analyzer calibration, setup and applications.

## Spectrum Analyzer Calibration

To use our spectrum analyzer in measuring applications we need to calibrate it and establish the largest spur-free dynamic range. There are two things to watch out for, overdriving the SA612A mixer and saturation of the sound card line-in amplifier. The SA612A mixer begins saturating at 20 mVRMS and the line-in saturates at about 1.0 VRMS in my case. So we don’t want to come close to or exceed these values. If we do, we will see a lot of distortion, as seen previously in Figure 8, which is a good tip-off that something is wrong. To get started set the program parameters as follows. Select the 24 kHz real time mode, check image removal, select medium FFT averaging and keep the spectrum peak away from the extreme lower and upper frequency ends where the sound card may roll off. This will produce a clean, stable and more accurate spectrum. A signal generator frequency in the range of 1 to 10 MHz works well.

In order to calibrate the analyzer we need to know the dBm or RMS voltage across R1, the 51 ohm resistor at the input of the mixer. A multimeter will not work here as most of them will not measure above a frequency of 10 kHz. A signal generator calibrated in dBm is the easiest method to use while a scope can also be used to measure the peak-to-peak voltage and then divide by 2.828 to convert to RMS. Then convert the RMS voltage to dBm using Equation 2.

Calibration is straightforward. With the program running, adjust the signal generator output to obtain a peak in the spectrum at the –10 dB mark of the display. The –10 dB line on the screen is used because that is the level used for the two-tone tests for IMD3 discussed later on. It also provides some overhead for larger signals. Now read the signal generator or scope to find the corresponding dBm value. Make a note of this value as it is the correction needed for calibration of the displayed dB scale. In my case the –10 dB mark on the screen corresponds to –36 dBm.

The perfectionist may wish to determine and/or modify the largest spur-free signal corresponding to the (–10 dB) mark on the analyzer screen. Here’s how its done.

First note that the mixer saturating input voltage of 20 mV corresponds to –21 dBm. We want to stay below this value to avoid distortion components. Even more so, we want to stay far below this point so as to not generate significant third order intermodulation distortion (IMD3). This means using a much lower mixer signal (for the –10 dB mark ) in order to get a good, spur free range but, hopefully, not so low as to be noise limited. This can be done by adjusting the gain of the op-amp by changing the values of R3 and R4.

The following has been found to be the case from IMD3 tests described later on. A value of –36 dBm, at the input of the mixer, produces a IMD3-free range of 82 dB on the screen from the (–10 dB level). Similarly, a –38 dBm value at the input of the mixer, produces a IMD3-free range of 84 dB. And so on. Clearly it is seen that a lower dBm level gives a larger spur-free range. But the lower you make this value the more the noise will affect the measurements at the low end below –90 dB on the screen.

Here’s how the adjustments are made. Assume, that we want –36 dBm for the –10 dB reference mark on the screen. This corresponds to about 3.5 mVRMS across the 51 ohm resistor R1 at the mixer input. Thus the 0 dB mark now corresponds to-26 dBm and a voltage of 11.5 mVRMS. The mixer has a conversion gain of about 17 dB and is followed by an amplifier of about 20 dB gain. Hence the total gain is 17 dB + 20 dB or 37 dB. This corresponds to a voltage gain of about 70.8. Hence the 11.5 mV signal should appear at the output of the operational amplifier at a level of about 814.2 mV. This is still below the saturation value of the sound card line input so there is some room for adjustment. This is done by adjusting the gain of the amplifier. The two resistors R3 and R4 are raised or lowered in value as needed to change the gain but must remain equal to each other to form a balanced input. Lowering the values, increases the gain. So this in effect lowers the dBm level corresponding to our -10 dB reference point. In my case, with the resistor values set at 20 k the –10 dB mark corresponds to –36 dBm. With the resistor values set at 15 k the –10dB mark corresponds to – 38 dBm. This is the value recommended in the circuit of Figure 4. To get a greater dynamic range you can choose 12K resistor values or lower. The down side is that you will be raising the noise floor. Remember that in operation always try to keep signals about 5 to 10 dB below the 0 dB line on the screen to allow for overhead and to provide a larger spur free range

Don’t expect extraordinary precision in the calibration as there are other factors at work. One of those is the FFT calculation. Since the FFT is using only a small part of the signal in the time domain (about 21.33 msec.), errors are introduced. This “time window”, sometimes referred to as a rectangular window, produces leakage of the signal into adjacent FFT frequency bins. To compensate, methods of windowing the data to reduce leakage have been developed. However none of them are perfect and, theoretically, cannot be since the leakage depends to a large degree on how well centered the signal is in the frequency bin. Nonetheless window functions can reduce the leakage significantly although there will still be small amplitude variations in the result. The different types of windows trade off selectivity, amplitude accuracy and noise floor. Several choices are available in the program. To learn more about FFT windows see the nice paper by Bores Signal Processing on the Web.5

Another factor is the amplitude variation of the LO. The direct digital synthesis (DDS) technology of the AD9851 uses a numerically controlled oscillator (NCO) to generate a digital sine wave. The digital sine wave is converted to analog form via an internal 10-bit high speed D/A converter causing the output amplitude to follow a sin(x)/x curve as shown in Figure 9. In this example the system clock is 125 MHz and the output frequency F0 is set to 20 MHz. In addition to the fundamental output F0, three aliases F1, F2 and F3 are shown and there are many more with increasing frequency that are not shown. The aliases are very prominent and of a relatively high energy level. In fact, the first alias F1 can equal the fundamental amplitude (when F0 = 1/2 system clock). This can be seen by noting that as the fundamental F0 moves to the right the first alias F1 moves to the left. A low-pass filter is generally placed at the output of the D/A converter to suppress the jitter-producing effects of non-harmonically related aliases and other spurious signals. Aliases are often thought of as unwanted by-products but in some cases, like this spectrum analyzer, can become the primary output signal. In the example shown, the F1 alias is at 105 MHz and quite useable. Not employing a filter and using aliases allows us to extend the range of the analyzer to almost the system clock frequency where its amplitude becomes too low to drive the mixer. However the calibration will be thrown off because the LO amplitude rolls off. Because of this and the above noted factors, the calibration is only good at the lower frequencies. Nevertheless, since many measurements and observations utilize only relative measurements and do not depend on absolute values, the spectrum analyzer is still useful in many situations even at the higher frequencies. The software makes note of the LO frequency and if it is above ½ system clock calculates the correct frequency for the alias being used.

There is one more complication. The DDS unit can operate from either a fixed crystal oscillator, like 125 MHz, or a lower frequency crystal oscillator and the internal 6x clock multiplying PLL loop of the AD9851. The phase noise performance of a PLL is determined by the multiplication ratio and the loop filter performance. Within the loop bandwidth of the multiplier, any noise that is present on the crystal clock will be increased in proportion to the multiplication value, here 6x. This inherent effect degrades performance of the DDS. The formula for signal degradation within the PLL loop bandwidth is given by dB = 20 log x, where x is the multiplication value of 6. Hence the signal degradation is about 15.5 dB. In my case when I compared a fixed 125 MHz crystal oscillator to a 30 MHz oscillator and 6x multiplier PLL, there was a significant increase in spurs for the PLL implementation. Therefore for best performance the fixed 125 MHz oscillator should be used. But since this is an experimenter’s project the program has provisions for using either method. When ordering the DDS module be sure to specify which frequency oscillator you want. Fortunately the PCB uses a socket and the oscillator can easily be swapped.

Now that we have checked out the analyzer and absorbed a little theory its time to look at some applications. This will also give us an opportunity to see examples of equipment testing, image removal, spectrum scanning, and real-time signal modulation.

## Signal Generators, Modulation, Filters And Other Devices

The first day the analyzer began working I decided to look at some signal generators in my home lab. What a surprise! To paraphrase the famous tale, “it was the best of times and the worst of times” as I discovered the idiosyncrasies of each one. To illustrate these cases, spectrums from various signal generators and a QRP rig are presented. We’ll also look at real time modulation, filter testing and more.

Figure 10 shows spectrums from two high quality signal generators. On the left is an analog tuned 10 to 510 MHz HP 8654A and on the right is a frequency synthesized10 kHz to 1 GHz Marconi 2022C, the latter recently being obtained from an internet auction site. As can be seen the 8654A has a much narrower spectrum all the way to the floor which indicates much lower noise. Unfortunately the 8654A drifts a lot and it takes over 15 minutes to settle down. On the other hand the Marconi unit has more noise but doesn’t drift much. The waterfall plots (not shown here) are excellent for tracking the drift and noise of signal generators. For these spectrums FFT averaging and image removal was employed.

Figure 11 shows spectrums from two lower frequency signal generators. On the left is a 5 MHz HP 3310A and on the right is a 50 MHz WaveTek Model 81. Compared to the previous generators, the noise floor for these units is much higher. It is particularly high for the Model 81 since it uses a diode network to convert the triangle signal to a sine wave. It also exhibits periodic step correction of the frequency during the warm-up which is easily seen on the analyzer but wasn’t noticed before. Also note the roll-off at the high frequency end due to the sound card

It’s easy to demonstrate AM and FM modulation using these signal generators as the analyzer can display the RF signals with only a small delay in one of the real time modes. Figure 12A show a typical AM signal with images and also the same signal, Figure 12B, with the image removal system engaged. The removal process adds a little time to the processing but is worth it in many cases as numerous spurs are also removed. Image removal is also demonstrated with an FM signal as shown in Figs. 13A and 13B. In this case the FM carrier,(normally at the center) is also missing since a modulation index of 2.4 is used. By the way, image removal in real time does not work as well with off-the-air signals as the sidebands are changing too fast. But wide band single scans do show AM carriers, video and FM signals fairly well.

QRP rigs can easily be tested with the analyzer. To illustrate, my old 2 W Ten-Tec PM2 QRP rig was used. It can be operated with a crystal or VFO control. I checked out both modes but will describe the VFO mode here. The PM2 was turned on, tuned up per the manual, and set to about 7.01 MHz. Because of the low power involved, the PM2 output was fed directly into the 50 ohm input of an attenuator connected to the analyzer. The waterfall plot was used to determine when the drift had stopped, which was about 15 minutes later. There was 60 dB of attenuation inserted in my setup to provide a peak at –10 dB as shown in Figure 14. This point corresponds to –36 dBm with my calibration. Hence, accounting for the input attenuator, the rig was putting out (60 – 36) dBm or 24 dBm (250 mW ) of power. With careful adjustment of the rig controls this could be increased to 350 mW. The harmonics were easily checked using the “X2” and “X3” LO buttons. The spectrum must be put on the left side of the screen before clicking these buttons so the harmonics will not fall off screen. Remember to click the button again to go back to the original spectrum. The second harmonic was found to be down 31.5 dB and the third harmonic down 45.3 dB from the carrier. This is often called dBc, or signal referenced to carrier.

Also shown is Figure 15, a 96 kHz stitched spectrum plot of the PM2. Note the dead spots or roll-off at 24 kHz intervals because of the sound card roll off. This is where the four 24 kHz spectrums are stitched together. Spectra is not valid at these points, so try to keep peaks away from those areas.

Insertion loss of devices can easily be measured. A bandpass filter (7.1 to 8.3 MHz), obtained on an auction site for $0.99, will be used to illustrate. First a signal generator and attenuator were connected to the analyzer and the attenuator adjusted for a convenient value of –10 dB at 7.6 MHz on the display. This is the reference value. Then the filter is inserted between the signal generator and attenuator and a new reading found. The difference in readings determines the insertion loss of the filter at this frequency. By varying the signal generator frequency, the 3 dB bandwidth points can be found. This requires coordinating the RF signal generator frequency and the analyzer’s LO to keep the spectrum on the screen and can be a little tricky. Commercial units use a tracking generator for this purpose. Another possibility is to use a noise generator to drive the filter instead of the signal generator. Unfortunately most low cost noise generators do not produce enough output to get a decent dynamic range unless an amplifier is used. A home-brew wide band noise generator using a Zener and followed by a LNA was used on the filter. It worked nicely and a frequency response plot was obtained. However, because of the limitation of the noise source, only about a 40 dB dynamic range was obtained and the response was in the lower part of the screen. Using a signal generator, instead of the noise source, produces a much larger dynamic range, so that method would be preferred. For this filter there was a 6 dB insertion loss and the bandwidth was spot on. Other components, such as a power splitter and combiner can be tested in a similar fashion as described next..

A power combiner was needed for the two tone tests described later on and so a Mini-Circuits ZFSC-2-2 was obtained on an auction site. The purpose of this device is to combine two signals into one while providing isolation of the two signal sources from each other. In this case it was a good idea to test the unit first using the same procedure used for the filter tested previously. For the ZFSC-2-2 there is typically 30 dB of isolation between ports 1 and 2 when the sum port S is terminated in 50 ohms, but there was only 3 dB loss measured! Upon opening the case it was seen that a wire had become unsoldered. Re-soldering the wire now produced the 30 dB isolation. Signal loss from port 1 or 2 to the sum port S can also be measured and was found to be 3.4 dB which is typical for this device.

## Two Tone Testing

Intermodulation distortion (IMD) is produced by virtually all amplifiers, mixers, filters and other devices due to the non-linear properties inherent in them. It occurs when two or more signals are applied to the device and describes a fairly common communications scenario; strong signals at neighboring frequencies introducing spurious signals into the current channel. Two tone testing for IMD is a time honored method for determining the strength of this interference and of the IMD products. A single, but important case, is covered here. However, for more information on the subject, you may wish to consult a source like an ARRL Handbook.

The frequency relationships in a two-tone IMD measurement are illustrated in Figure 16. Frequencies f1 and f2 are applied to the DUT and third order products (IMD3) at (2f1–f2) and (2f2–f1) are generated. The relative magnitude of these products defines the IMD. Although many IMD products are created in a test like this, the IMD3 products usually cause the most problems since they fall into the bandpass region.

Figure 17 shows the test setup for two tone testing. It consists of two signal generators (Marconi and HP in my case) for f1 and f2, a power combiner (ZFSC-2-2), the device under test (DUT), attenuator and the spectrum analyzer. Other signal generators could be used here and there are many good ones available on the surplus market. Homebrew crystal oscillators and power combiners, sometimes called hybrids, are also another option.

The first item that needs testing is the mixer in the spectrum analyzer since its distortion products will be added to those of any other device tested. The SA612 is not the best choice for a mixer here as it generates more IMD at lower input levels than a balanced diode mixer. It was chosen, however, because it is low cost, sensitive, consumes little power, requires only modest LO power, and has a large +17 dB conversion gain requiring little post-amplification. Experimenters may want to consider the possibility of designing a new front end for the analyzer if IMD turns out to be a problem for some of their tests. However I have found the SA612 can be used here with a good dynamic range that belies its reputation for IMD3. This is true for several reasons.

The SA612 and other mixers are usually measured for IMD3 at large offset frequencies up to 1 MHz, (60 kHz for the SA612). This is because they are often used in receiver front ends where significant off-frequency interference occurs. That is normally not the case with our spectrum analyzer in the 24 kHz mode. Using lower offsets of 2 to 6 kHz improves the IMD3 performance of the SA612 considerably. Lower offsets are not a drawback since many tests on amplifiers and SSB gear often use 2 to 3 kHz offsets. Another thing that can lower IMD3 is to raise the power supply voltage of the SA612 and op-amp to 6.5 volts. This is safely below the 8 volt maximum. And finally, because the SA612 is so sensitive, it can be operated at lower input levels where IMD3 is lower.

Here the IMD3 mixer tests are done with the f1 and f2 peaks set to the –10 dB mark. This is done for two reasons. One is that the peak power is 6 dB above the spectrum peak and could overdrive the analyzer if set higher. And secondly it provides improved spur free dynamic range (SFDR) for the IMD3 measurement. The IMD3 tests on the mixer are done using the setup of Figure 17, but with the DUT replaced by a short. Therefore most of the IMD3 showing on the screen will be due to the SA612 mixer in the analyzer. Some IMD3 will also be due to the signal generators interacting but that interaction is difficult to reduce without using a more complicated setup. Also some will be due to IMD3 products from the sound card but hopefully they are well below the mixer IMD3 levels.

Figs. 18A and 18B show two-tone results for this analyzer’s mixer for two different calibration levels. Note that with the calibration reference (–10 dB point) set to –36 dBm the SFDR is 82 dB. And with it set to –38 dBm it is 86 dB. This is enough to get down to the noise floor at the bottom of the plot. If desired the range can be extended farther as described in the calibration section. By the way, you can always tell if that is the IMD3 component on the screen by changing the attenuator by 1 dB and seeing the IMD3 level change by 3 dB.

To illustrate a practical case, a Ramsey Model PR2 RF broadband (1 MHz to1 GHz) pre- amplifier was used as the DUT in Figure 17 and tested for IMD3. The amplifier in the PR2 is a UPC 1678 MMIC with a gain of 23 dB and a 1 dB compression point at 18 dBm referred to the output. Using two-tones near10 MHz with about 4 kHz spacing produced the plot shown in Figure 19. The tones are at –10 dB (–36 dBm with my calibration) and IMD3 is at –68 dB (–94 dBm). The attenuator is set to 38 dB indicating that the MMIC is producing +2 dBm at its output and an IMD3 spur of –56 dBm. For comparison, the data sheet for the MMIC shows a plot of IMD3 versus SSB output power. It indicates that +2 dBm output power corresponds to an IMD3 of about –53 dBm. Since IMD3 for the MMIC varies with factors like temperature and power supply voltage, the value measured here is reasonably close. The gain of the PR2 amplifier can also be measured using the procedure described previously for filters.

In case you want to check more powerful amplifiers you may need to insert an amplifier before the “amplifier under test” as many signal generators do not have enough power output to drive them directly. Make sure the driving amplifier has IMD well below the one being tested to avoid clouding the results.

AM/SSB/CW Receiver and Radio Spectrum Scans

A filter and amplifier should be added to the front end if we want to use the spectrum analyzer as a receiver. Otherwise its wideband capability and weak IMD3 performance will produce noisy results. A good antenna, bandpass filter and the PR2 amplifier have worked well here for all ham bands up through 10M. If you use the lowest RF bandwidth (24 kHz) without image removal you will be able to hear the signals, corresponding to the lower end of the spectrum on the screen, through your sound card. You can hear AM and SSB by zero beating the signal. Its also great fun tuning in to the CW parts of the band and observing the individual dots and dashes (for the slower code speeds), on the waterfall plot. During a CW contest, dozens of signals can be seen at once, cascading down the waterfall. An example is shown in Figure 20.

Its also fun to observe larger portions of the radio spectrum on the analyzer. The single scan modes ranging from 192 kHz to 12.2 MHz are best. Since the signals will be weak, a wide-band amplifier such as the PR2 mentioned earlier should be used at the input. An example spectrum of part of the TV and FM band starting at 82 MHz is shown in Fig 21. Video from Channel 6 is seen on the left side near 83 MHz and numerous FM stations, most notably WUWM, the campus station is seen at 89.7 MHz. In this case the analyzer was using alias frequencies. When aliases are employed it may be helpful to utilize a filter at the input of the analyzer to remove those signals below ½ the clock frequency.

As one last example, the lowly AM broadcast band was checked one day and an unusual signal at 1130 kHz. was found. See Figure 22 for the spectrum. Can you guess what those strange looking rectangular sidebands are? Yes, its HD AM radio (from Ibiquity Digital) and those are digital sidebands next to the AM ones.6 They are about 30 dB down from the carrier and within FCC limits. Its hard to believe that even the AM radio band now has digital signals!

## Final Analysis

This spectrum analyzer can prove helpful in a variety of situations where a more expensive unit is not available. Even though its design is simple, it can be employed for routine equipment testing, component checkout, educational purposes and as a receiver and signal generator. It has some unusual features that may prove interesting to experimenters such as the use of alias frequencies, image removal and spectrum scanning. Compared to commercial spectrum analyzers of a few years ago, its noise floor is quite good and its narrow band sweep speed is much faster.

Drawbacks include marginal IMD performance, limited sound card frequency response, LO amplitude roll-off and very slow wideband sweep speed. Some of these things can be improved with a more complicated design and the possible use of a 96 kHz sound card. Nevertheless, as it stands, it is a useful instrument.

One desirable addition would be a RF preselector filter, as used with many spectrum analyzers, which would facilitate distortion-free measurements in the presence of large signals at the input. Filtering helps by significantly attenuating undesired out-of-band signals, reducing energy from high amplitude, short duration pulses, and minimizing the total number of individual signals present at the analyzer input. It would consist of tunable low pass and high pass filters or several band-pass filters that could be switched in and out. This would improve on my bothersome method of using individual filter modules.

It was instructive and a lot of fun designing, building and testing the analyzer. When you get your unit working, please let me know about it. In any case, let the world be your spectrum and start analyzing it now.

Notes

5 www.bores.com.

6www.ibiquity.com/hd_radio.

Figure 9 — Typical sin(x)/x output power curve for AD9851. (File SAFig9.doc)

Figure 10 — Spectrums of HP 8654A and Marconi 2022C high frequency signal generators. (Editor please choose either SAFig10HPMarconia.gif or SAFig10HPMarconib.gif enhanced)

Figure 11 — Spectrums of HP 3310A and WavTek Model 81 low frequency signal generators. (Editor please choose either SAFig11HPwavea.gif or SAFig11HPwaveb.gif enhanced)

Figure 12A — AM modulation without image removed. (Editor please choose either SAFig12AAMmodulationNIRa.gif or SAFig12AAMmodulationNIRb.gif enhanced)

Figure 12B — AM modulation with image removed. (Editor please choose either SAFig12BAMmodulationIRa.gif or SAFig12BAMmodulationIRb.gif enhanced)

Figure 13A — FM modulation without image removed. (Editor please choose either SAFig13AFMmodulationNIRa.gif or SAFig13AFMmodulationNIRb.gif enhanced)

Figure 13B — FM modulation with image removed. (Editor please choose either SAFig13BFMmodulationIRa.gif or SAFig13BFMmodulationIRb.gif enhanced)

Figure 14 — Older Ten-Tec PM2 QRP rig spectrum at 24 kHz bandwidth. (Editor please choose either SAFig14_24kHzTenTecPM2a.gif or SAFig14_24kHzTenTecPM2b.gif enhanced)

Figure 15 — Older Ten-Tec PM2 QRP rig showing stitched spectrum at 96 kHz bandwidth. (Editor please choose either SAFig15_96kHzTenTecPM2a.gif or SAFig15_96kHzTenTecPM2b.gif enhanced)

Figure 16 — Frequency relations for two tone IMD tests. (File SAFig16.doc)

Figure 17 — Setup for two tone IMD tests. (File SAFig17.doc)

Figure 18A — SA612 mixer two tone spectrum with the calibration reference set to –36 dBm . (Editor please choose either SAFig18A20KMixertwoToneTesta.gif or SAFig18A20KMixertwoToneTestb.gif enhanced)

Figure 18B — SA612 mixer two tone spectrum with the calibration reference set to –38 dBm. (Editor please choose either SAFig18A15KMixertwoToneTesta.gif or SAFig18A15KMixertwoToneTestb.gif enhanced)

Figure 19 — Ramsey PR2 preamplifier two tone spectrum. (Editor please choose either SAFig19RamseyPR2TwoTonea.gif or SAFig19RamseyPR2TwoToneb.gif enhanced)

Figure 20 — Spectrum and waterfall plot for CW activity. (Editor please choose either SAFig20SpectrumWaterfalla.tif or SAFig20SpectrumWaterfallb.tif enhanced)

Figure 21 — Scan of TV and FM band starting at 82 MHz — (Editor please choose either SAFig21TVandFMbanda.gif or SAFig21TVandFMbandb.gif enhanced)

Figure 22 — Spectrum of AM radio station WISN (1130 kHz) using HD radio. (Editor please choose either SAFig22AMHDRadioa.gif or SAFig22AMHDRadiob.gif enhanced)